Security screening – Management Licence holder’s obligations

Management licence condition M1766 Document reference: HHSD/22/434859

OFFICIAL

# Introduction

The Victorian Radiation Act 2005 (the Act) has the objective of protecting the health and safety of persons and the environment from the harmful effects of radiation. The Department of Health (Department) administers this legislation.

The Act seeks to fulfil this objective by establishing a licensing framework to regulate the conduct of radiation practices and the use of radiation sources.

Any person who conducts a radiation practice must hold a management licence (unless exempted from that requirement). The management licence holder must comply with every condition of their licence.

# Scope

This document describes the obligations of management licence holders where condition M1766 has been imposed in respect of an authority to possess a whole-body transmission security scanner. The requirements in this document apply to the possession of the following apparatus for the purpose of security screening;

- Whole body transmission and backscatter security scanner - Section 1; and

- Industrial radiography source (incorporating an X-ray unit) - Section 2.

# Requirements

## Whole body transmission and backscatter security scanner

### 1.1 Security scanner compliance

    - For a Backscatter Security Scanner, the licence holder must ensure that the security scanner complies with the System and Manufacturing Requirements for a General-use System specified in the American National Standard ANSI/HPS N43. 17-2009;

    - For a Whole Body Transmission Security Scanner, the licence holder must ensure that the security scanner complies with the System and Manufacturing Requirements for a Limited-Use System specified in the American National Standard ANSI/HPS N43. 17-2009;

### 1.2  Warning notices

1.2.1	The licence holder must ensure that:

- illustrated notices requesting that a person informs staff before the security screening that they might be pregnant are prominently displayed in the area of the security scanner.

- each access point for a member of the public into the area of the security scanner has a warning sign to indicate that the rooms contains an X-ray apparatus.

### 1.3  Limitation of radiation exposure

1.3.1	The licence holder must ensure that radiation doses to occupationally exposed persons and members of the public are kept as low as reasonably achievable, economic and social factors being taken into account.

1.3.2	The licence holder must ensure that working conditions of an occupationally exposed female are adapted when the female declares that she is pregnant to ensure that the radiation exposure of the embryo or foetus is below the prescribed dose limit for a member of the public as specified in Table B of Schedule 4 of the Radiation Regulations 2017.

1.3.3	The licence holder must ensure that:

- the reference effective dose for the scanner is determined in accordance with Section 6.1.3, “Determination of the Reference Effective Dose,” in the American National Standard ANSI/HPS N43. 17-2009;

- a record is kept of the reference effective dose and the number of scans an individual has undergone in any 12-month period; and

- administrative controls in the form of documented procedures are implemented to ensure that the effective dose to any individual scanned is limited to 0.70 mSv (700 μSv) in any 12-month period.

### 1.4  Personal radiation monitoring

1.4.1	The licence holder must ensure that:

- a personal radiation monitoring device is provided to each occupationally exposed person;

- a record is kept of the radiation doses received by each occupationally exposed person; and

- work practices are investigated and reviewed if an occupationally exposed person receives an effective dose in excess of 0.25 mSv in any three-month period.

### 1.5  Radiation Management Plan

1.5.1	The licence holder must develop, document, implement, and review at intervals not exceeding twelve months a Radiation Management Plan that incorporates the following components:

the maximum radiation dose per scan for a person exposed to the security scanner;

a system to ensure that the maximum annual radiation dose that can be received by a person who is subjected to security screening with the security scanner does not exceed 0.70 mSv;

the type of records kept in relation to the scanned person;

the retention of radiation dose records for persons exposed to radiation from the security scanner;

the licensing requirements of the radiation regulatory authority;

the actions necessary to manage a radiation incident, including reporting and investigation of the radiation incidents; and

the record of security scanner calibration results.

### 1.6  Shielding

1.6.1	The licence holder must maintain radiation shielding in the doors, walls, floor and ceiling of the room in which the radiation source is installed and any shielding for operators which is necessary to ensure that no person receives a radiation dose in excess of the relevant radiation protection limit specified in the Radiation Regulations 2017.

## Industrial radiography source (incorporating an X-ray unit)

### 2.1  Radiation Management Plan

2.1.1	The licence holder must ensure that a Radiation Management Plan is developed, documented, implemented and regularly reviewed to ensure safety in all licensed activities, including the following:

- roles and responsibilities;

- work practices, including the provision of PPE;

- radiation monitoring requirements, including details of how the availability or accessibility requirements for the monitoring equipment are to be achieved;

- control and management of incidents;

- storage and transport, and disposal of radioactive material (if applicable);

- accountability and records; and

- any other requirement that may have a bearing on safety.

# Further information

Email:		Radiation.Safety@health.vic.gov.au

Website:	https://www.health.vic.gov.au/radiation/security-screening

To receive this document in another format, email Radiation Team <Radiation.Safety@health.vic.gov.au>. Authorised and published by the Victorian Government, 1 Treasury Place, Melbourne. © State of Victoria, Australia, Department of Health, May 2026. Available at Radiation website < https://www.health.vic.gov.au/publications/security-screening-management-licence-holders-obligations >